## B. Change Log
Appendix B: Change Log
| Requester information |  | Change Proposal Information | [10 empty] | IMDRF Decision
|  |  | Identification of code / term for which proposal is made | [5 empty] | Proposal of change
No. | Date submitted (DD/MM/YYYY) | Submitter (organisation name) | Terminology (Annex A, B, C, D, E, F, G) | Version of Annex | Code | Term | Location in the hierarchy | Definition | Category of change (Select: Add, delete, modify) | Description of change (e.g. "modification of definition …") | Rationale for change (e.g. "the change is necessary to accommodate a new type of device…") | Impact on other existing terms | Example of an incident which would be coded using the proposed term | Outcome of change request (ACCEPTED or REJECTED) | Justification | New code if applicable | Date Published (DD/MM/YYYY)
1 | 2020-08-28 | MHRA | E | 2.1 | E24 | Others | L1 |  | Modify | Change status to "Not selectable" | All other L1 codes are marked as not selectablein annex E, because they are used only for grouping. The same should apply for E24, since on its own the term "other" doesn't tell us anything, and one of the three more specific terms must be selected. | None | N/A | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
2 | 31/08/2020 | PMDA | E | 2.1 | E0304 | Hypervolemia | Level 2 | Too much fluid in the blood. | Modify | Change the definition to either "Too much fluid in the blood vessel"or "An abnormally high circulating blood volume." (the latter aligns with the definition for Hypovolemia(E0305): An abnormally low circulating blood volume.) | Since blood is a type of fluid, too much fluid in blood is unclear. We suggest being more specific. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed, using the latter definition. | N/A | 02/03/2021
3 | 31/08/2020 | PMDA | E | 2.1 | E0508 | Low Pulmonary Arterial Wedge Pressure | Level 2 | The measurement of the mean left pulmonary atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is low. | Modify | Modify the definition as follows: The measurement of the mean left pulmonary atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is low. | The "left pulmonary atrial pressure" does not make anatomical sense. It also aligns better with the definition for High Pulmonary Arterial Wedge Pressure (E0507): The measurement of the mean left atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is high. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
4 | 31/08/2020 | PMDA | E | 2.1 | E082901 | Acanthameba Keratitis | Level 3 | Keratitis due to infection by acanthamoba. | Modify | Modify the term and definition to use the following spelling: Acanthamoeba | The spelling should be consistent in the term/definition. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
5 | 31/08/2020 | PMDA | E | 2.1 | E1101 | Biliary Cirrhosis | Level 2 | Cirrhosis of the liver caused either by destruction of the intrahepatic bile ducts (primary biliary cirrhosis) or blockage of the extrahepatic bile ducts (secondary biliary cirrhosis). | Modify | Modify the definition as follows: Cirrhosis of the liver caused either by destruction of the intrahepatic bile ducts (primary biliary cholangitis) or blockage of the extrahepatic bile ducts (secondary biliary cholangitis). | To more accurately reflect the diagnosis of the disease (cirrhosis does not develop until later stages of the disease), the WHO approved the change from primary biliary cirrhosis to primary biliary cholangitis in 2016, and ICD-11 codes also uses this term. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
6 | 31/08/2020 | PMDA | E | 2.1 | E0105 | Cerebral Ventriculomeglia | Level 2 | Abnormal enlargement of the cerebral ventricles. | Modify | Modify the term to "Cerebral Ventriculomegaly" | This is more commonly used (the MedDRA term is also this term). | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
7 | 31/08/2020 | PMDA | E | 2.1 | E1625 | Non-union Bone Fracture | Level 2 | Broken bone that doesn't heal. | Modify | Modify the term to "Nonunion Bone Fracture" | This aligns with MedDRA and ICD. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
8 | 31/08/2020 | PMDA | E | 2.1 | E1904 | Subclinical Infection | Level 2 | Infection associated with no detectable symptoms but caused by microorganisms capable of producing easily recognizable diseases. | Modify | Modify the definition as follows: Infection associated with no detectable signs or symptoms but caused by microorganisms capable of producing easily recognizable diseases. | No detectable signs should also be included as a condition for subclinical infections. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
9 | 01/09/2020 | PMDA | E | 2.1 | E2106 | External Prosthetic Device Pain | Level 2 | Pain that is associated with an external ariticial replacement device. | Modify | Modify the definition as follows: "Pain that is associated with an external artificial replacement device." | Editorial change. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
10 | 31/08/2021 | PMDA | A | 2.1 | A010202 | Loss of Osseointegration | Level 3 | Problem associated with weakened integration of the device at the bone-implant interface due to loss of fibrous and/or bony tissue and leading to compromised anchorage of the device. i.e. 'Loosening/Lysis\"\" | Modify | Remove extra symbols. | Editorial change. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. *Note: The commenter appears to have reviewed an older version of the terminology. This correction was already made in Edition 4.1 published on 27 July 2020. | N/A | 27/07/2020
11 | 31/08/2021 | PMDA | A | 2.1 | A0901 | Audible Prompt/Feedback Problem | Level 2 | Problem with any deviation from the documented specifications of the device that relate to audible feedback. e.g. voice prompts or beeps, but not safety related alarms which are covered under \"\"Protective Measures Problem\"\". | Modify | Remove extra symbols. | Editorial change. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. *Note: The commenter appears to have reviewed an older version of the terminology. This correction was already made in Edition 4.1 published on 27 July 2020. | N/A | 27/07/2020
12 | 31/08/2021 | PMDA | A | 2.1 | A090504 | Unexpected/Unintended Radiation Output | Level 3 | Device-emitted radiation when it was not supposed to. This applies to devices which are intended to emit radiation, and the radiation being emitted from the correct part of the device, but at an incorrect time. Use \"\"radiation leak\"\" if the device emits radiation which should never have been emitted, or from a location from which it should never be emitted. | Modify | Remove extra symbols. | Editorial change. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. *Note: The commenter appears to have reviewed an older version of the terminology. This correction was already made in Edition 4.1 published on 27 July 2020. | N/A | 27/07/2020
13 | 31/08/2021 | PMDA | A | 2.1 | A140301 | Inability to Auto-Fill | Level 3 | Complete failure to fill as part of an automated process. For insufficient filling use \"\"Short Fill\"\". For excessive filling use \"\"Overfill\"\". For inconsistent filling use \"\"Volume Accuracy Problem\"\". | Modify | Remove extra symbols. | Editorial change. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. *Note: The commenter appears to have reviewed an older version of the terminology. This correction was already made in Edition 4.1 published on 27 July 2020. | N/A | 27/07/2020
14 | 31/08/2021 | PMDA | A | 2.1 | A140302 | Overfill | Level 3 | Excessive filling of a device. For complete failure to fill use \"\"inability to auto-fill\"\". For insufficient filling use \"\"Short Fill\"\". For inconsistent filling - use \"\"Volume Accuracy Problem\"\". | Modify | Remove extra symbols. | Editorial change. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. *Note: The commenter appears to have reviewed an older version of the terminology. This correction was already made in Edition 4.1 published on 27 July 2020. | N/A | 27/07/2020
15 | 31/08/2021 | PMDA | A | 2.1 | A140303 | Short Fill | Level 3 | Insufficient filling of a device. For complete failure to fill use \"\"inability to auto-fill\"\". For insufficient filling use \"\"short fill\"\". For inconsistent filling - use \"\"Volume Accuracy Problem\"\". | Modify | Remove extra symbols. | Editorial change. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. *Note: The commenter appears to have reviewed an older version of the terminology. This correction was already made in Edition 4.1 published on 27 July 2020. | N/A | 27/07/2020
16 | 31/08/2021 | PMDA | A | 2.1 | A140304 | Volume Accuracy Problem | Level 3 | Inconsistent filling of a device. This describes a problem which is observed to vary between overfilling and under filling, and may be intermittent. Use \"\"Overfill\"\" or \"\"Short Fill\"\" if problem is consistent. | Modify | Remove extra symbols. | Editorial change. | None. | NA (Not a new term). | ACCEPTED | Agreed as proposed. *Note: The commenter appears to have reviewed an older version of the terminology. This correction was already made in Edition 4.1 published on 27 July 2020. | N/A | 27/07/2020
17 | Sep. 4, 2020 | AdvaMed | D | 2.1 | D14 | No Problem Detected | 1 | The device complaint or problem cannot be confirmed. | Modify | Modification of definition to: "The device complaint or problem cannot be confirmed or was confirmed through the investigation to be unrelated to the device." | Currently there is no code that captures after an investigation that the adverse event was unrelated to the manfacturer's device. The current definition only seems to capture if there is insufficient information, such that it cannot be confirmed. | None | Manufacturer receives a report that seems to suggest their device is involved in an event, however during the investigation information is received confirming there was no device related issue. | ACCEPTED | IMDRF accepted this Change Request with modifications: This comment indicates a misunderstanding of the intended meaning of D14. If it can be confirmed that an event was not related to the device then D10 should be used. The definition for D10 is being clarified in response to another comment. D14 should be used when the reported problem canot be reproduced, but this is better captured as a finding in Annex C. Change C19 definition to clarify meaning. Term: No Device Problem Found New Definition : Reported problem cannot be reproduced during investigation. | N/A | 02/03/2021
18 | Sep. 4, 2020 | AdvaMed | D | [5 empty] | Add | Addition of code for "Not Manufacturers Device" | Currently there is no code that captures after an investigation reveals the event did not involve the manufacturer's device. | None | Manufacturer receives a report that seems to suggest their device is involved in an event, however during the investigation information is received demonstrating a label or product identification number showing another manufacturer's device. | ACCEPTED | IMDRF disagreed with the proposed term, but agreed to add a note to the definition of B16. B16 (Device Not Manufactured by Reporting Manufacturer) covers the key point in this case and doesn't need to be repeated in anexes C and D. In this case the most appopriate terms are C20 (No Findings Available) and D14 (No Problem Detected). B16 Modified Definition: Further information was obtained which established that the manufacturer of the device involved was not the one to which it was initially attributed. Use in conjunction with C20 and D14. | N/A | 02/03/2021
19 | Sep. 4, 2020 | AdvaMed | D | [5 empty] | Add | Addition of code for “Report Retraction" | Currently there is no code that captures a retraction of the adverse event report itself. | None | Manufacturer receives a report that seems to suggest their device is involved in an event, however during the investigation information is received demonstrating the information no longer meets the criteria for reportability. | REJECTED | "Report Retraction" is not an Investigation Conclusion, but it may be useful to capture these in a consistent way. Some jurisdictions have an option to retract reports or submit a final report indicating the event doesn't meet reporting criteria. IMDRF will consider this in further work on terminology. | N/A | N/A
20 | Sep. 4, 2020 | AdvaMed | E | 1 |  |  |  |  | Add | Addition of "No Reportable Adverse Event, No Clinical Signs, Symptoms, Conditions" | Currently there is no code that captures when there is no reportable adverse event such that there is no reportable "Health Effects - Clinical Signs and Symptoms or Conditions" Code. E2403 assumes that there was a reportable adverse event but the patient had no clinical signs or symptoms. | None | In an abundance of caution a manufacturer initially reports an event and codes a "Health Effects - Clinical Signs and Symptoms or Conditions", however through the investigation it is determined no reportable adverse event occurred, there is no code that captures there are no longer reportable patient health effects. | REJECTED | IMDRF does not agree with the satement that "E2403 assumes that there was a reportable adverse event." If there was a clinical effect which is "not reportable" then there is more value in coding it correctly. Some jurisdictions have an option to retract reports or submit a final report indicating the event doesn't meet reporting criteria. IMDRF will consider this in further work on terminology. | N/A | N/A
21 | Sep. 4, 2020 | AdvaMed | F | 1 |  |  |  |  | Add | Addition of "No Reportable Adverse Event, No Health Impact" | Currently there is no code that captures when there is no reportable adverse event such that there is no reportable "Health Effects - Health Impact" Code. F26 and F27 both assume that there was a reportable adverse event but there was either no harm or no patient involvement. | None | In an abundance of caution a manufacturer initially reports an event and codes a "Health Effects - Health Impact", however through the investigation it is determined no reportable adverse event occurred, there is no code that captures there are no longer reportable patient health impact. | REJECTED | IMDRF does not agree with the satement "F26 and F27 both assume that there was a reportable adverse event". The fact that a report was submitted suggests that there was a reportable event. If the Health Impact was not known initially then use F24 "Insufficient Information". There is more value in coding why an event is not reportable. When a follow-up or final report is submitted using F26 or F27 then it will be clear. Some jurisdictions have an option to retract reports or submit a final report indicating the event doesn't meet reporting criteria. IMDRF will consider this in further work on terminology. | N/A | N/A
22 | Sep. 4, 2020 | AdvaMed | B |  |  | Sample Requested But Not Provided | Top Level | A sample (patient, survey, etc.) was required to adequately investigate the issue; the sample was requested but not provided. | Add |  | Investigations of IVD issues that appear to be due to interferences or other sample-related issues require a sample for investigation in order to determine a root cause. We have a term for "Device Not Returned." A similar code is necessary to adequately describe the efforts of the manufacturer to perform a complete investigation when a required sample is not sent to the manufacturer. | None | The customer alleges TSH results on a sample are very different on 2 different manufacturers' instruments. An interfering factor to a component of one of the reagents is suspected. A sample is required for investigation, but the customer could not provide the sample. | ACCEPTED | IMDRF accepted this Change Request with modifications: ADD New Term Annex B Term: Specimen Requested But Not Provided Definition: A patient specimen was required to adequately investigate the issue and was requested but not provided. | B23 | 02/03/2021
23 | Sep. 4, 2020 | AdvaMed | A |  | A0908xx | Quality Control Invalid | L3 | Quality control test produces an invalid result. | Add | Addition of codes to differentiate between patient and QC results | <not necessarily required for adverse event reporting in the US, but required for IVDR trending in Europe> Applies to IVD - currently there is no code that captures QC results. Failed QC might be an accompanying issue to adverse events (patient results) or be standalone issue. | better differentiation vs. Other A0908xx |  | REJECTED | This is not a description of a device problem, so does not fit in Annex A. There are already terms which cover the actual device problem. Problems identified during QC testing are not in scope for manufacture reporting in some jurisdictions, but the results of testing may indicate a patient safety issue. IMDRF will consider these types of events in further work on terminology. | N/A | N/A
24 | Sep. 4, 2020 | AdvaMed | A |  | A0908xx | Quality Control Out of Range | L3 | Quality control result out of range (e.g. after calibation, maintenance). | Add | Addition of codes to differentiate between patient and QC results | <not necessarily required for adverse event reporting in the US, but required for IVRD trending in Europe> Applies to IVD - currently there is no code that captures QC results. Failed QC might be an accompanying issue to adverse events (patient results) or be standalone issue. | better differentiation vs. Other A0908xx | On an IVD analyzer, after daily maintenance (user action), QC results for one or more analyzers are out of range, indicating an issue with the device. | REJECTED | This is not a description of a device problem, so does not fit in Annex A. There are already terms which cover the actual device problem. Problems identified during QC testing are not in scope for manufacture reporting in some jurisdictions, but the results of testing may indicate a patient safety issue. IMDRF will consider these types of events in further work on terminology. | N/A | N/A
25 | Sep. 4, 2020 | AdvaMed | A |  | A0908xx | External Quality Control/Panel | L3 | External quality control/panel does not meet expected results for external survey (e.g. EQA, CLIA, CAP, Ring trial, EQAS, PHE, SKLM). | Add | Addition of codes to differentiate between patient and EQA results | Applies to IVD - currently there is no code that captures EQA results. EQA samples run like a real patient sample might be subject to reportability in certain jurisdictions. | better differentiation vs. Other A0908xx | A CAP sample run on an IVD system returns result grossly differing from peer group, indicating an issue with the device. | REJECTED | This is not a description of a device problem, so does not fit in Annex A. There are already terms which cover the actual device problem. Problems identified during testing such as EQA are not in scope for manufacture reporting in some jurisdictions, but the results of testing may indicate a patient safety issue. IMDRF will consider these types of events in further work on terminology. | N/A | N/A
26 | Sep. 4, 2020 | AdvaMed | D |  | D01xx | Cybersecurity | L2 | Problems traced to inadequate protection of computers, servers, mobile devices, electronic systems, networks, and data from malicious attacks, damage, or unauthorized access. | Add | No specific codes for cybersecurity | Currently, we are using D0105 and D1107, but these are still a bit far off from cybersecurity | None | One customer’s cobas IT 1000 server has been infected with the ransomware called “Blammo ransomware”. The Blammo ransomware has been run and it has infected the customer network. The cobas IT 1000 server files were encrypted and not usable. | ACCEPTED | IMDRF accepted this Change Request with modifications. Added a New L2 Term under D01 (Cause Traced to Device Design) and definition to Annex D New Term: Inadequate Cybersecurity New Definition: Problems traced to inadequate protection of computers, servers, mobile devices, electronic systems, networks, programs, and data from malicious attacks, damage, or unauthorized access. | D0107 | 02/03/2021
27 | Sep. 4, 2020 | AdvaMed | D |  | Dxx/D01xx | Unintended Software Error | L1/L2 | Problems traced to an error, flaw or fault in a computer program or system that causes it to produce an incorrect or unexpected result, or to behave in unintended ways | Add | SW error is not a Product Design Error (which is more Hardware) | currently, the option we have is to select "Design problem", which is clearly incorrect. A SW error is not always related to product design | None | There is no visual indication in POCT Management> Blood gas overview when a cobas b 123 is in “System stop”. | ACCEPTED | IMDRF accepted this Change Request with modifications. Added New terms and definitions under Annex D. New Term (L1): Cause traced to software coding. New Definition: Problems traced to an error, flaw or fault in a computer program or system that causes it to produce an incorrect or unexpected result, or to behave in unintended ways. New term (L2 under D01): Inadequate Software Design. New Definition: Cause traced to inadequate or inappropriate design of the software component of a medical device, or Software as a Medical Device. | D18 D0108 | 02/03/2021
28 | Sep. 4, 2020 | AdvaMed | D | 4.1 | D10 | Cause Cannot be Traced to Device | 1 | The adverse event that occurred is not attributable to a device. | Modify | Modify Definition: The reported incident was confirmed to have occurred, but it was not device related. Cause is unknown. | Clarity for the use of D10, D14, D15 codes is needed or the codes will be used interchangeably. | Yes, definition changes for D10, D14 and D15 | It was reported (and confirmed via imaging) that the patient experienced a pericarial effusion. Additional investigation determined that this occurred prior to the device being implanted. The investigation was furthermore unable to determine if it was related to the procedure or a pre-existing patient condition. | ACCEPTED | IMDRF accepted this Change Request with modifications: This comment indicates a misunderstanding of the intended meaning of D10. D10 should be used when the cause was confirmed to be something other than the reported device. Annex C should be used to indicate problems which were confirmed. If despite these findings the cause is unclear then the appropriate term is D15 - Cause Not Established. Modified D10 Term: Caused traced to non-device related factors Modified Definition: The adverse event that occurred is confirmed to have been caused by something other than the device or usage. | N/A | 02/03/2021
29 | Sep. 4, 2020 | AdvaMed | D | 4.1 | D14 | No Problem Detected | 1 | The device complaint or problem cannot be confirmed. | Modify | Modify Definition: The reported incident was refuted and not related to device, user, procedure or the patient. | Clarity for the use of D10, D14, D15 codes is needed or the codes will be used interchangeably. | Yes, definition changes for D10, D14 and D15 | Use for all cases where the investigation determined that the incident did not occur as reported, therefore a cause code is not applicable. | ACCEPTED | IMDRF received two contradictory comments on D14. D14 should be used when the reported problem canot be reproduced, but this is better captured as a finding in Annex C. This term would still apply if it were confirmed that the reported incident didn't happen, although that is not the only meaning of this term. Change C19 definition to clarify meaning. Term: No Device Problem Found New Definition : Reported Problem cannot be reproduced during investigation. | N/A | 02/03/2021
30 | Sep. 4, 2020 | AdvaMed | D | 4.1 | D15 | Cause Not Established | 1 | The investigation findings do not lead to a clear conclusion about the cause of the reported adverse event. | Modify | Modify Definition: The reported incident was device related but the cause could not be determined. | Clarity for the use of D10, D14, D15 codes is needed or the codes will be used interchangeably. | Yes, definition changes for D10, D14 and D15 | Use for all cases where there was clearly a device issue that caused the reported incident but the investigation could not identify the cause (e.g. Design, Manufacturing, Component). | ACCEPTED | IMDRF disagreed with the rationale for this Change Request, but agreed to add L2 term under D15. This comment seems to indicate a misunderstanding of the various annexes. Reported device problems should be coded using A, confirmed device problems should be coded using C, and the root cause of the problem should be coded using D. D15 is intended to be used when there is no clarity at all about the cause, whether or no the reported problem could be confirmed during investiagtion. However, IMDRF agrees that it would be helpful to have a term to indicate that an event is linked to the device but the root cause remains unclear. Add New term and definition under D15 "Cause not established." L2 New Term: Cause Linked to Device but Unable to Trace More Specifically New Definition: Problems traced to the device, but the investigation could not identify the actual root cause (e.g. Design, Manufacturing, Component). | D1501 | 02/03/2021
31 | Sep. 4, 2020 | AdvaMed | D | 4.1 | D15 [Note added by IMDRF: This should be D12] | Known Inherent Risk of Device | 1 | Reported adverse event known and documented in the labeling (including both short or long term known complications or adverse reactions). | Delete | Remove this code unless it can be made more clear when it should be used. | This is not a cause but rather a catch-all that will allow a manufacturer to not capture a cause as long as the adverse event is listed in lableling. Many potential adverse events are listed in labeling but they still have a cause. This code should be deleted or a very straightforward definition provided to make it clear when it should be used in absence of a true cause. | Yes, definition changes for D10, D14 and D15 | Pericardial Effusion is listed in the physician user manual as a potential adverse event that may occur. However, if the investigation determines that the clinician used excessive force when implanting the lead the manufacturer should use a Cause Traced to User Code. | ACCEPTED | IMDRF accepted this Change Request and added a clarification as to when this term should be used: Modified D12 definition : Reported adverse event known and documented in the labeling and all reasonable mitigation steps have been taken. (including both short or long term known complications or adverse reactions). | N/A | 02/03/2021
32 | Sep. 4, 2020 | AdvaMed | C | 4 | ALL | ALL | ALL | ALL | Modify | Add clarity to the IMDRF/AE WG/N43FINAL:2020 (Edition 4) guidance document | Cause coding (annex D) is focused on identifying the probable cause whether device, patient, user or environmental. However, annex C is device centric. The guidance document should be updated to provide instruction on what "Investigation Findings" codes should be added when the cause is not device related. | No | Health care provided reported that the patient's lead became dislodged 1 week post implant procedure, likely as a result of the patient throwing a football. Investigation Findings code C19 (No Device Problem Found) applied. | ACCEPTED | IMDRF accepted this Change Request with modifications: IMDRF accepted that user problems cannot be coded using annex C, although maintenance and environment problems can be, and added a new L1 term to Annex C as indicated below. However, the example given in the Change Request should not be coded using C19 if a problem was found (lead dislodged), even if the cause was related to exceptional use. Annex C is focused on what problems were found and not what they were caused by. ADD Annex C L1 New Term: Usage Problem Identified Definition : Problems that occured related to the actions of a healthcare professional, patient, or other device user. | C23 | 02/03/2021
33 | 2020-10-22 | MHRA | E | 2.1 |  | Bone Cement Implantation Syndrome | L2 | A rare and potentially fatal perioperative complication of cemented bone surgery, characterised by hypotension, hypoxia, cardiac dysrhythmias, and in severe cases cardiac arrest. | Add | Add new term to enable monitoring of a rare but serious complication associated with bone cement. | There is a rare but serious event which is of particular interest. It can currently only be identified by coding the symptoms. | None | "Patient was having a cemented prosthesis inserted when they had a cardiorespiratory arrest immediately after the insertion of the cement. Unresponsive to vasopressors and fluid replacement. This appears to be a case of Bone Cement Implantation Syndrome." | ACCEPTED | Agreed as proposed. | E2120 | 02/03/2021
34 | 2020-10-27 | US FDA | E | 2.1 |  | Multiple Organ Failure | L2 Under Metabolism and Nutrition (corresponds to MedDRA placement) | Complete impairment of two or more organs or organ systems | Add | Add new term to enable monitoring of multiple organ failure related to medical deivces | This code was previously used frequently related to ventricular assist devices. The term "multiple organ failure" is what is typically used by clinicians and provided to the manufacturer without providing more specific information. | None | IT WAS REPORTED THAT THE PATIENT BECAME SEVERELY VASODILATED AND VOLUME WAS DEPLETED POST IMPLANT. THE PATIENT REQUIRED METHYLENE BLUE AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT AND WAS GIVEN INTRAVENOUS (IV) FLUIDS. THE SITE REPORTED THAT THE PATIENT WAS HYPOTENSIVE, HAD AN ELEVATED LACTATE LEVEL, AND RENAL FAILURE. PATIENT SUBSEQUENTLY EXPIRED. THE CAUSE OF DEATH WAS DUE TO MULTI SYSTEM ORGAN FAILURE. | ACCEPTED | IMDRF accepted this Change Request with modifications. E23 Generalized Disorders L2 New Term: Multiple Organ Dysfunction Syndrome New Definition: Complete impairment of two or more organs or organ systems. Also known as multiple organ failure. MedDRA code: 10077361 | E2342 | 02/03/2021
35 | 21/01/2021 | MHRA/PMDA | G | 1.1 | G07 | Others | L1 | Terms not yet classified. If these terms get moved to other categories, this category can be deleted. (Previously the device category.) | Modify | Change definition to "Others." Add status and status description indicating that this term is not selectable. | The current definition is an old working note used during drafting of the annex. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
36 | 15/02/2021 | PMDA | A | 2.1 | A0908 | Incorrect, Inadequate or Imprecise Resultor Readings | L2 | Problem associated with a nonconforming end result, data, or test results provided by the device to its performance specifications. | Modify | Change term to "Incorrect, Inadequate or Imprecise Result or Readings" (added a space between Result and or) | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
37 | 15/02/2021 | PMDA | A | 2.1 | A15 | Activation, Positioning or SeparationProblem | L1 | Problem associated with any deviations from the documented specifications of the device that relate to the sequence of events for activation, positioning or separation of device. Note: Deployment is synonymous with activation. | Modify | Add a space between Separation and Problem. "Activation, Positioning or Separation Problem" | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
38 | 15/02/2021 | PMDA | C | 2.1 | NA | Annex Instructions | NA | Terms/Codes in this annex should be used to to answer the question What were the results of the investigation? | Modify | Remove an extra "to." "Terms/Codes in this annex should be used to to answer the question What were the results of the investigation?" | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
39 | 15/02/2021 | PMDA | D | 2.1 | NA | Annex Instructions | NA | Terms/Codes in this annex should be used to to answer the question What is the conclusion regarding the root cause? (based on the results noted in the Investigation Findings and the way they were obtained noted in the Type of Investigation.) | Modify | Remove an extra "to." Terms/Codes in this annex should be used to to answer the question What is the conclusion regarding the root cause? (based on the results noted in the Investigation Findings and the way they were obtained noted in the Type of Investigation.) | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
40 | 15/02/2021 | PMDA | D | 2.1 | D1107 | Sabotage or Intentional Misuse | L2 | Problems occurred because of wilful damage or reckless misuse. This term indicates that there was an apparent intent to cause damage or injury. | Modify | Problems occurred because of willful damage or reckless misuse. This term indicates that there was an apparent intent to cause damage or injury. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
41 | 15/02/2021 | PMDA | E | 2.1 | NA | Annex Instructions | NA | These terms allow capturing of the patient signs and symptoms observed related to the medical device adverse event through observational language without using diagnostic specifics. These terms are largely based on a subset of MedDRA terms. This annex is organized along organ systems as well as physiological problems. The Level 1 term is mainly for organizational purposes and and is expected to be too high level for coding purposes. | Modify | Remove an extra "and." These terms allow capturing of the patient signs and symptoms observed related to the medical device adverse event through observational language without using diagnostic specifics. These terms are largely based on a subset of MedDRA terms. This annex is organized along organ systems as well as physiological problems. The Level 1 term is mainly for organizational purposes and and is expected to be too high level for coding purposes. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
42 | 15/02/2021 | PMDA | E | 2.1 | E010904 | Status Epilepticus | L3 | A life-threatening condition caractarized by a single proplonged seizures or a series of seizures without intervening full recovery of conciousness. | Modify | A life-threatening condition charactarised by a single proplonged seizures or a series of seizures without intervening full recovery of consciousness. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
43 | 15/02/2021 | PMDA | E | 2.1 | E0115 | Encephalopathy | L2 | A functional and/or structural disorder of the brain that is acquired or congenital and charac terised by mental and neurological symtoms. | Modify | A functional and/or structural disorder of the brain that is acquired or congenital and characterised by mental and neurological symptoms. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
44 | 15/02/2021 | PMDA | E | 2.1 | E0135 | Taste Disorder | L2 | Definition:conditions characterized by an alteration in taste function or perception | Modify | Definition:conditions characterized by an alteration in taste function or perception | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
45 | 15/02/2021 | PMDA | E | 2.1 | E0710 | Choking | L2 | Interference with respritation by ostruction of the larynx or trachea. | Modify | Interference with respiration by obstruction of the larynx or trachea. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
46 | 15/02/2021 | PMDA | E | 2.1 | E0749 | Strangulation | L2 | Interference with respiration by compression of the larynx, laryngopharynx or trachea,as well as of the carotid arteria or jugular veins. | Modify | Interference with respiration by compression of the larynx, laryngopharynx or trachea, as well as of the carotid arteria or jugular veins. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
47 | 15/02/2021 | PMDA | E | 2.1 | E0835 | Toxic Anterior Segment Syndrome (TASS) | L2 | An acute postoperative inflammatory reaction in which a noninfectious substance enters theanterior segmentand inducestoxicdamage to the intraocular tissues. | Modify | Added several spaces between words: An acute postoperative inflammatory reaction in which a noninfectious substance enters the anterior segment and induces toxic damage to the intraocular tissues. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
48 | 15/02/2021 | PMDA | E | 2.1 | E0843 | Zonular Dehiscence | L2 | Rupture of the fibrrous strands connecting the ciliary body with the crysaline lens of the eye (usually at the time of cataract surgery). | Modify | Rupture of the fibrrous strands connecting the ciliary body with the crystalline lens of the eye (usually at the time of cataract surgery). | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
49 | 15/02/2021 | PMDA | E | 2.1 | E1305 | Renal Impairment | L2 | An acute or chronic condition where the kidnies are not functioning normally. | Modify | An acute or chronic condition where the kidneys are not functioning normally. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
50 | 15/02/2021 | PMDA | E | 2.1 | E1409 | Nipple Sensation Changes | L2 | Excessive sensitivity or changes in sensitivty in the nipple(s) | Modify | Excessive sensitivity or changes in sensitivity in the nipple(s) | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
51 | 15/02/2021 | PMDA | E | 2.1 | E1412 | Rectovaginal Fistula | L2 | Abnormal connetion between rectum and vagina. | Modify | Abnormal connection between rectum and vagina. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
52 | 15/02/2021 | PMDA | E | 2.1 | E1905 | Transmissible Spongiform Encephalopathy(TSE) | L2 | A group of rare degenerative brain disorders associated with orions for example Creutzfeldt-Jakob disease (CJD) or Bovine Spongiform Encephalopathy (BSE). | Modify | A group of rare degenerative brain disorders associated with prions for example Creutzfeldt-Jakob disease (CJD) or Bovine Spongiform Encephalopathy (BSE). | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
53 | 15/02/2021 | PMDA | E | 2.1 | E2119 | Unintended Radiation Exposure | L2 | The unintentional exposure to radiation of a person, or a part of a persons body. This might include exposure of a healtchare professional, or the exposure of the wrong part of the body of a patient who was intended to receive therapeutic or diagnostic radiation. This includes both ionizing and non ionizing radiation. | Modify | The unintentional exposure to radiation of a person, or a part of a persons body. This might include exposure of a healthcare professional, or the exposure of the wrong part of the body of a patient who was intended to receive therapeutic or diagnostic radiation. This includes both ionizing and non ionizing radiation. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
54 | 15/02/2021 | PMDA | E | 2.1 | E231201 | Extreme Exhaustion | L3 | Extreme fatique; the inability to repond to stimuli. | Modify | Extreme fatigue; the inability to respond to stimuli. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
55 | 15/02/2021 | PMDA | E | 2.1 | E231501 | Purulent Discharge | L3 | A flow or draining of pus or exudate from an oriface or wound. | Modify | A flow or draining of pus or exudate from an orifice or wound. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
56 | 15/02/2021 | PMDA | E | 2.1 | E2336 | Shock | L2 | A life-threatening condition that requires immediate medical intervention.It is characterized by reduced blood flow that may result in damage of multiple organs. | Modify | A life-threatening condition that requires immediate medical intervention. It is characterized by reduced blood flow that may result in damage of multiple organs. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
57 | 15/02/2021 | PMDA | F | 2.1 | F14 | Prolonged Episode of Care | L1 | Clinically significant extended treatment time or procedureal time. Not to be used for prolonged hospitalization or prolonged surgery. | Modify | Clinically significant extended treatment time or procedureal time. Not to be used for prolonged hospitalization or prolonged surgery. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
58 | 15/02/2021 | PMDA | F | 2.1 | F2204 | IVD Testing | L2 | Pateint required additional or unexpected IVD testing. | Modify | Patient required additional or unexpected IVD testing. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
59 | 15/02/2021 | PMDA | F | 2.1 | F2305 | Radiation Therapy | L2 | Patient required addtional radiation therapy. | Modify | Patient required additional radiation therapy. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
60 | 15/02/2021 | PMDA | F | 2.1 | F25 | Unanticipated Adverse Device Effect | L1 | Complication occuring in the patient that had not previously been identified or expected for this device. | Modify | Complication occurring in the patient that had not previously been identified or expected for this device. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
61 | 15/02/2021 | PMDA | F | 2.1 | F26 | No Health Consequences or Impact | L1 | No apparent harm ocurred in relation to the adverse event. | Modify | No apparent harm occurred in relation to the adverse event. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
62 | 15/02/2021 | PMDA | G | 1.1 | G02011 | Device Reader | L2 | A piece of hardware used to read the memory and the properites of a device. | Modify | A piece of hardware used to read the memory and the properties of a device. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
63 | 15/02/2021 | PMDA | G | 1.1 | G0201501 | Electrode | L3 | A small piece of metal or other conductive substance this is used to take an electric current to or from a source of power, a piece of equipment.Not to be used for patient connection. | Modify | A small piece of metal or other conductive substance this is used to take an electric current to or from a source of power, a piece of equipment. Not to be used for patient connection. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
64 | 15/02/2021 | PMDA | G | 1.1 | G02020 | Heater | L2 | A piece of equipment used to raise the termperature of air, gas or water or an object. | Modify | A piece of equipment used to raise the termperature of air, gas or water or an object. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
65 | 15/02/2021 | PMDA | G | 1.1 | G03006 | Cuvette | L2 | A clear contatiner designed to interface with an optical sensor in order to obtain an optical measurement of a contained substance. | Modify | A clear contatiner designed to interface with an optical sensor in order to obtain an optical measurement of a contained substance. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
66 | 15/02/2021 | PMDA | G | 1.1 | G0405202 | Bolt | L3 | A cylindirical connector element which may have a thread/nut connection to form a fastener. | Modify | A cylindirical connector element which may have a thread/nut connection to form a fastener. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
67 | 15/02/2021 | PMDA | G | 1.1 | G04095 | Pad | L2 | Typically a soft, cushion-like material used to prevent injury or damage to a person or device. Can also be made of an absorbant material to absorb liquid. | Modify | Typically a soft, cushion-like material used to prevent injury or damage to a person or device. Can also be made of an absorbent material to absorb liquid. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
68 | 15/02/2021 | PMDA | G | 1.1 | G04130 | Tool | L2 | A component that is delivered with a device to support its assemby or operation, and is not part of the device. Use only in the case there are no other terms to describe the component. | Modify | A component that is delivered with a device to support its assembly or operation, and is not part of the device. Use only in the case there are no other terms to describe the component. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
69 | 15/02/2021 | PMDA | G | 1.1 | G0600102 | Alarm, Visual | L3 | A component designed to signal the occurence of a particular event in a way that can be seen. | Modify | A component designed to signal the occurrence of a particular event in a way that can be seen. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021
70 | 15/02/2021 | PMDA | G | 1.1 | G0600601 | Needle Stick Prevention Mechanism | L3 | A mechanism intergrated into a device to prevent needle stick injuries. | Modify | A mechanism intergrated into a device to prevent needle stick injuries. | Editorial change. | None | NA (Not a new term). | ACCEPTED | Agreed as proposed. | N/A | 02/03/2021

## Lists (hidden)
Annex | Change | Outcome
A | Add | ACCEPTED
B | Delete | REJECTED
C | Modify
D
E
F
G